Metallurgical Microscope

MJ22

User Guide

Congratulations and Thank You for your purchase of The Microscope

This precision instrument has been designed to require a minimum of optical and mechanical maintenance. Its excellent design assures years of high quality, reliable service. We recommend you read this entire manual carefully before beginning to use the instrument.

### WARNING

PLEASE DISCONNECT PLUG FROM MAINS SOCKET BEFORE REMOVING THE BOTTOM PLATE, OPENING THE LAMP REPLACEMENT DOOR, OR REMOVING THE LAMP HOUSE

### WARNING

THE POWER CORD PROVIDED WITH THE EQUIPMENT HAS A GROUNDED PLUG. ALWAYS USE THE POWER CORD WITH A PROPERLY GROUNDED WALL OUTLET.

DO NOT EXPOSE THE INSTRUMENT TO HIGH TEMPERATURES OR HUMIDITY. AVOID USING THE INSTRUMENT IN EXTREMELY DUSTY LOCATIONS

OPERATING TEMPERATURE 5 DEGREE C. TO 35 DEGREE C.

OPERATING HUMIDITY: 20% TO 80% @ 25 DEGREE C.

CAUTION:	NEVER IMMERSE THE INSTRUMENT IN WATER OR SOLVENT

CAUTION:	DO NOT PUT ANY FOREIGN OBJECTS IN THE FRAME OR INTO ANY MOVING MECHANICAL PARTS

Table of Contents

Getting Started	3

Components	4

Setup Instructions	5

Basic Operation	6

Changing the Incident (upper) Light Bulb	12

Changing the Transmitted (lower) Light Bulb	12

Preventative Maintenance	13

Getting Started

This reference guide is written base on whole series, it covers some capable optional attachments and functions, you may read the similar operation what you need.

Components

Setup Instructions

Remove all parts from their packing materials and retain the packaging in the event you need to transport the product.

Loosen the thumbscrew and install the Optical Body with Vertical Illumination Tube to the frame by removing its dust cap and inserting the dovetailed flange on the underside of the optical body into the frame. Secure optical body with thumbscrew. DO NOT RELEASE THE OPTICAL BODY FORM YOUR GRASP UNTIL IT IS FIRMLY

SECURED TO THE FRAME.

3.	Install the objective into the nosepiece

- Remove the dust caps from the eyetubes in the Optical Body and insert

both eyepieces into the Optical Body.

- Install the lamp house

- Connect the power cord to the instrument and to a suitable power supply.

Basic Operation

- Illumination controls

a) The power switch to the illuminator and the brightness control are located on the base. The transmitted / incident illumination toggle switch 1(Figure 2B) is located on the right side. The electrical system is fuse protected and the is fuse holder, located in the base of the microscope.

b) Turn on the light with the Power switch 1 (Figure 2A), see Figure 2. If the light does not appear to be ON, check the brightness control 2( Figure 1A) to see if it’s on a sufficiently low setting. Also check the transmitted / incident switch is set on the transmitted setting.

- Focusing Controls

a) Focusing adjustment is accomplished by using the large coarse adjustment knobs located comfortably on each side of the frame. Fine adjustment is accomplished using the smaller knobs located on the same focus shaft. This coaxial arrangement allows for easy, precise adjustment without drift or discomfort.

b) Focus Control Turning either of the coarse focus control knobs will raise or lower the stage. The smallest graduation on the fine adjust knob index scale is 2μm of vertical.

c) Focus Tension Adjustment The tension of the coarse focus is adjustable and preset at the factory for ease of use. If you wish to adjust the coarse focus tension, first locate the tension adjustment ring on the left coarse adjustment knob. It is located between the frame and right coarse adjustment knob. Turning the ring toward the rear of the microscope increases the tension, and toward the front of the microscope loosens it. Tension is too high if you experience physical discomfort

d) Pre-focusing or Focus Stop Control This lever is located on the right side of the frame, between the coarse focus knob and the frame. Use of this feature will insure that the longer objectives don’t contact the stage or specimen when using the microscope. Its use also simplifies focusing. After focusing on the specimen with the coarse adjustment, rotation of the lever toward the front of the microscope will set and upper limit on the coarse adjustment movement. After changing specimens or objectives, focusing is easily accomplished by rotating the coarse adjustment knob to reach the pre-focused position, then making fine adjustments with the fine adjustment knob. Focusing movement with the fine adjustment isn’t affected by using the pre-focusing lever

- Interpupillary and Diopter Adjustments

- Diopter Adjustment Proper correction for individual vision is accomplished via the diopter adjustment located at the left eyepiece, see Figure 4. Using the 40X objective, close your left eye and bring and image into focus with your right eye only. Once the image is well focused, close your right eye and check the focus with your left eye. If the image is out of focus, make fine adjustments with the diopter adjustment ring to correct for your vision.

- Proper interpupillary distance, or the distance between eyepieces, is crucial to the comfort of the user. Adjusting the interpupillary distance is accomplished through a “folding” action of the Seidentopf design optical head, see Figure 4, allowing for quick and easy adjustment

- Mechanical Stage Controls

The mechanical stage, see Figure 5, with    Ultra-Low position control knobs for ease of use, consists of three parts:

a) The smaller knob for X-Axis or transverse    adjustment.

b) The larger knob for Y-Axis or longitudinal adjustment.

- Condenser Components – are located below the stage as follows, see Figure 6: Abbe condenser; its Mounting Ring with lock screw and two large centering screws at either side for adjusting condenser/iris assembly in the optical center of the light path; the adjustment knob for raising and lowering the condenser lens. The substage lamp house includes a field iris diaphragm for Kohler illumination and a recess for installation of a filter, if desired.

- Condenser Alignment Adjustments to the    substage condensing system, see Figure 7, are crucial for proper illumination and performance. There are three basic adjustments to be made: centering, vertical focusing and aperture adjustment.

  - Centering–The condenser must be centered in the light path to ensure proper light control. One simple method for centering follows:

i) Rotate the objective nosepiece until        the 5X objective is in the light path.

ii) While looking into the microscope,    close the field diaphragm to the smallest setting using the knurled ring at the top of the light exit window assembly.

iii) Closing the iris in this manner will reduce the field so that a small white circle is visible within a black field, see Figure 8A. This white circle is the light that is passing through the iris and it should be centered in the black field as in Figure 8b. If it is not centered, move it to the center using the centering screws.

iv) By opening the iris until the white circle almost fills the entire field, see Figure 8C.

  - Vertical Focusing–The condenser can be raised and lowered with the substage adjustment knob in order to focus the light for optical illumination.

  - Aperture Adjustment–The contrast can be adjusted with the iris diaphragm adjustment lever on the condenser assembly. Aperture adjustments are made to induce contrast into a specimen, not to adjust light intensity.

- Vertical Illumination Alignment, See Figure 9.

  - To center the incident or Vertical lamp put a small piece of paper and place it on the stage. Take out one objective form nosepiece and rotate the nosepiece until this objective opening hole is in position over the paper.

  - Turn the incident(upper) illumination on and adjust the lamp intensity control to establish sufficient light.

  - This paper will allow for focusing of the lamp filament on the paper.

Refer to Figure 2 for this step. reduce the aperture iris diaphragm, using the adjustment lever and rotate the lamp focusing control on the lamp housing until a sharp image of the lamp filament is attained.

  - Refer to Figure 10 for this step. If the filament image is not centered in the overpass hole, loosen the lamp set screw and move the bulb socket with the lamp centering knob to obtain a uniform image. Retighten the set screw.

  - Reopen the aperture iris diaphragm.

  - Reinstall the objective again

  - The Iris diaphragm nearest the illuminator house in the vertical illuminator tube may be closed slightly to enhance contrast. The iris diaphragm nearest the illuminator house should be centered for Kohler illumination.

  - There is a field diaphragm near the trinocular body for Kohler Illumination.

  - The Filter wheel contains a green, yellow, blue, and neutral density filters, as well as a clear opening.

  - The Polarizing Analyzer is used in conjunction with the Polarized filter cube, selected by pushing in the blue filter rod (Figure 11). The Analyzer may also be placed over the lower light source on the base (Figure 7) and used with the Polarized filter cube for transmitted light work.

  - The white push rod to the upper left of the trinocular head will send 100% of the image light to the binocular eyepieces or to the top port.

Changing the Incident (upper) Light Bulb

- Turn off and unplug the microscope.

- Loosen the lamp set screw, Figure 12, unscrew the lamp centering control lever and remove the entire lamp socket from the lamp house.

- The halogen bulb can be removed by grasping the bulb and pulling it firmly from its fixture. Take care not to touch the lamp with bare fingers, as the lamp will be hot. Insert the new halogen bulb into the same fixture. When installing the new bulb, be careful not to touch the glass with your fingers. The new lamp should be supplied in a plastic protective envelope. If not, use a tissue or other medium to grasp the bulb. This will prevent contamination from your hand from reducing the bulb’s intensity and life.

- Reinsert the lamp socket into the lamp house and retighten the setscrew. Make certain that the new lamp is placed squarely in the socket to insure that is will center properly in the field of view. You may need to re–center the lamp as previously described.

Changing the Transmitted (lower) Light Bulb

Disconnect the power plug and wait to be sure the bulb is cool. Locate the hinged door on the bottom of the frame. The lamp door retaining screw is to the front of the frame, and the lamp centering screw is located in the middle of the lamp door. Open the lamp door using the retaining screw as shown in Figure 13. pull the old lamp straight out of the socket. Always use a clean cloth or lens tissue to prevent grease or oil getting on the glass of the new

bulb. Insert the lamp pins straight into the socket. Close the lamp door and center the lamp in the field of view using the lamp positioning screw at the center of the lamp door and pulling it firmly, see Figure 13

Preventative Maintenance

- Cleaning frame and stage

Disconnect the plug from mains socket before cleaning. Clean the frame and stage with a soft cloth moistened with a mild detergent solution. Be sure the instrument is dry before using.

Cleaning optical parts

Microscope eyepieces and objectives are coated.

They should not be wiped while dry as dirt or dust may scratch the coating. It is best to remove parts from the frame prior to cleaning. Always blow loose dust away first. Use cotton swabs or lens tissue moistened with a lens cleaner or a small amount of alcohol, then wipe the surface clean with a good quality lens tissue. Solvents such as Xylene should NOT be used as cleaner.

DO NOT DISASSEMBLE OBJECTIVE LENSES.

# Text boxes

15

Overpass Hole

4

3

2

1

Figure 3

1. Fine Focus 2. Focus Tension 3. Up Stop 4. Coarse Focus

Figure 10

Filament

Figure 8

C

B

A

Figure 11

1

2

3

4

5

6

7

8

9

10

11

12

13

14

16

17

18

19

20

21

22

23

24

1

25

26

27

Figure 1

1. Eyepiece (2)  2. Trinocular Body  3. Polarizer  4. Field Iris Diaphragm  5. Field Iris Diaphragm Centering Screws  6. Filter Wheel  7. Aperture Iris Diaphragm  8. Screw of Lamp House  9. Focusing Lens Lever  10. Lamp House  11. Lamp Adjustment Lever  12. Locking Screw  13. Coarse Focusing Knob  14. Fine Focusing Knob  15. Plug for Vertical Illumination  16. Transmitted / Incident Illumination Toggle Switch  17. Lengthwise Knob 18. Cross Knob  19. Collector 20. Field Diaphragm 21. Abbe Condenser and Iris Diaphragm with Filter Tray  22. Condenser Centering Screws 23. Abbe Condenser Locking Screw 24. Stage 25. Objectives  26. Tube Holding Screw  27. Analyzer Lever

2

A

1. Power switch 2. Brightness control

Figure 2

B

1. Transmitted / Incident illumination toggle switch

1

Diopter Adjustment Ring

Figure 4

Y-Axis

X-Axis

Figure 5

1

2

3

4

5

7

6

Figure 6

1. Stage  2. Abbe Condenser Centering Screws  3. Abbe Condenser Lock Screw  4. Iris Diaphragm  5. Abbe Condenser 6. Substage Lamp House with Field Iris Diaphragm  7. Filter Holding Tray

Substage Adjustment Knob

Figure 7

1

11

10

1. Camera Port  2. Push Rod to Direct Image to Binocular Eyepieces or Camera Port (White) 3. Blue Rod for Normal Upper Light  4. Green Rod for Analyzer Filter

2

9

3

4

8

7

5

6

Figure 9

1. Trinocular Body  2. Field Iris Diaphragm  3. Filter Wheel  4. Aperture Iris Diaphragm 5. Lamp Adjustment Lever  6. Locking Screw  7. Focusing Lens Lever  8. Screw of Lamp House  9. Field Iris Diaphragm Centering Screws  10. Polarizer  11.Analyzer lever

3

4

1

2

1

2

3

4

Figure 12

1. Lamp Adjustment Lever 2. Lamp Set Screw  3. Bulb  4. Lamp Socket

Figure 13

1. Bulb 2. Lamp Door Retaining Screw  3. Lamp Centering Screw

1

3

2