Polarizing Microscope

User Manual

Micro-shot Technology Limited

www.m-shot.com

Thanks for your purchase of the Microscope MP30

This precision instrument needs your careful maintenance. To ensure its service year, we recommend you read this entire manual carefully before using.

WARNING

PLEASE DISCONNECT PLUG FROM MAINS SOCKET BEFORE REMOVING THE BOTTOM PLATE, OPENING THE LAMP REPLACEMENT DOOR, OR REMOVING THE LAMP HOUSE

WARNING

DO NOT EXPOSE THE INSTRUMENT TO HIGH TEMPERATURES OR HUMIDITY.

AVOID USING THE INSTRUMENT IN EXTREMELY DUSTY LOCATIONS.

OPERATING TEMPERATURE 5℃ TO 35℃

OPERATING HUMIDITY: 20% TO 80% (25℃)

CAUTION:	NEVER IMMERSE THE INSTRUMENT IN WATER OR SOLVENT

CAUTION:	DO NOT PUT ANY FOREIGN OBJECTS IN THE FRAME OR INTO ANY MOVING MECHANICAL PARTS

1

Content

Getting Started	3

Components 	3

Setup Instructions	5

Basic Operation	6

Changing lamp and fuse	10

Maintenance	10

2

Getting started

This user manual is written on basis of MP30 series microscope, it contains all components and attachments it may have, please read this user manual by your needs.

Components

1	2	3

4

5

6

7

8

Figure 1

- Connector lock-screw 2. Body lock-screw 3. Bertrand lens centering screw 4. Slide holder 5. Condenser (with aperture diaphragm 6. Polarizer 7. Condenser (With field diaphragm and filter set 8. Brightness control knob 9. Power 10. Fine adjustment hand wheel 11. Tension adjustment hand wheel 12. Condenser lift hand wheel  13. Stage centering screw

3

1	2	3	4	5

|  | 6 |
| --- | --- |
| 13 |  |
| 12 | 7 |
| 11 |  |
| 10 |  |

9	8

Figure 2

1.Eyepiece  2. Trinocular head  3. Bertrand lens  4. Analyzer  5. Extend neck lock screw 6. Limit stopper hand wheel 7. Course and fine system 8. Power line 9. Power seat 10. Objective 11. Condenser  12. Compensator  13. Extend neck

4

Setup instruction

- Remove packages of all parts, you can keep the packages for needs in the future.

- Loose body lock screw, rotate trinocular head and face to instrument, then tight lock screw.

- Push Bertrand lens, polarizer and analyzer to matched position.

- Fix 40X objective to nosepiece which does not have adjustment screw, then fix other objectives one by one.

- Take off eyepiece dust cover, then insert in eyepieces.

- Connect power line .

5

| Basic operation | Power control button |
| --- | --- |
| 1. | Illumination |  |
|  | 1) | Power control button and brightness control wheel is on the bottom of the microscope. The whole electric system is protected by the fuse, fuse seat is inside the power seat. |  |
|  | 2) | Open power and make illumination working, as Figure 3. If not bright enough, please check whether brightness control is under too low position. | Figure 3 |
| 2. | Focusing |  |

- Coarse and fine focusing system

is adjusted by 1 2 3 4 coarse hand

wheel on both sides of microscope

body. Fine focusing is achieved

by coaxial.

| 2) | By rotate any side of hand wheel you can focusing and lift stage with minimum moving division 2μm. |  |  |  |
| --- | --- | --- | --- | --- |
| 3) | Before shipment, the coarse focusing has been settled to a suitable tension position. While you can find tension hand wheel at the right side focusing hand wheel to rotate.Too tight is not comfortable to use instrument. | Figure 4 Fine focusing hand wheel Tension hand wheel Limit stopper than wheel Coarse focusing hand wheel |

- Preset limit stopper hand wheel can make sure when using shot working distance objective, it will not forced to stage or slide. Steps: when using low times objective, adjust coarse focusing to get clear field of view with sample, rotate forward to yourself it can set limit position. After changing to other objectives, then you can easily rotate coarse fusing system to suitable place and adjust fine focusing system. Limit stopper does not work to find focusing.

6

3. Diopter and InterpupillaryAdjustment

- Diopter Adjustment Proper correction for individual vision is accomplished via the diopter adjustment located at the left eyepiece, see figure 4. Use the 40X magnification objective, bring an image into focus with your right eye only. Once the image is well focused, observe with left eye, make fine adjustments with the diopter adjustment ring to correct for your vision.

- Proper interpupillary distance or the distance between eyepieces, is crucial to the comfort of the user. Adjusting the interpupillary distance is accomplished through a “folding” action of the optical head, at figure 5, allowing quick and easy adjustment.

- Rotate stage as Figure 6

1) Put the specimen on the satge, observe it with division eyepiece and 40 X objectives.

- Fix a target point in the field, shfit the specimen until the target and the ceter coincide. Refer to figure 7A.

- Rotary stage, if the stage hasn’t beed center-fixed, the target would run in a circle track. Refer to figure 7B.

- Adjust the ceter- fixed screw, make the cetre of the track close to the field centre, then the ceter-fixed is finished. Refer to figure 7C.

- If other objectives are not on the center, as figure 7B, then adjust the objective center-fixed screw until the objective is in the center of the optical axis.

6) Tighten the locking screw to prevent the stage from spinning

Diopter adjust ring

Figure 5

1	2	3	4

Figure 6

1Stage centering screw  2. 40X objective

| 5. polarizer observation, figure 10 | 1 | 2 |
| --- | --- | --- |

- Push the analyzer the the operation position,

leave the anlyzer and the polarizer both on “0”

(figure 10, to trafer the observation mode to

quadrature. It means to make the polarizer on

the west and east while the analyzer on the

south and north.

- Insert λ, λ/4 or quartzes compensator

if needed.

- The observe/camera switch pushrod is

for the shift between observing and picking

up. The camera adopt 100% trasmission to

meet a higher need.

Attention: You may need to readjust the

polarizer whe doing a trasmission observation

if the collector has been adjusted before, refer

to figure 10. Push the analyzer to the light

path and leave both the polarizer and analyzzer

on “0”. Losen the collector locking screw, spin

the collector until the field being observed reaches

its darkest, and then relock the screw.

6. conoscopic observation, figure 10.

Push the Bertrand lens to the working position, leave the machine on the conoscopic observation mode. Adjust the Bertrand lens center-fixing screw to remove the image to the field center.

9

Replacing the Light Lamp and the fuse

- Take off power,disconnect power line, make sure lamp has cooled down, loose lamp lock screw and take out lamp seat from microscope seat.

- Pull the old lamp straight out of the socket. Insert the new one into the lamp socket. When installing the new bulb, be careful not to touch the glass with your fingers. The new lamp should be supplied in a plastic protective envelope. If not, use a tissue or other medium to grasp the bulb. This will prevent contamination from your hand from reducing the bulb’s intensity and life. Switch on the power, adjustment brightness of illumination, check normal or not.

Maintenance

1. Cleaning frame and stage

Disconnect the plug from mains socket before cleaning. Clean the frame and stage with a soft cloth moistened with a mild detergent solution. Be sure the instrument is dry before using.

- Cleaning optical parts

Microscope eyepieces and objectives are coated.

They should not be wiped while dry as dirt or dust may scratch the coating. It is best to remove parts from the frame prior to cleaning. Always blow loose dust away first. Use cotton swabs or lens tissue moistened with a lens cleaner or a small amount of alcohol, then wipe the surface clean with a good quality lens tissue. Solvents such as Xylene should NOT be used as cleaner.

3.  Cleaning 100X oil immersion lens

The immersion oil should be removed from the lens at the end of each workday using cotton swabs or lens tissue moistened with a lens cleaner or a small amount of  alcohol.

DO NOT DISASSEMBLY COMPONENTS BY YOURSELF.

# Text boxes

MP30 Series

13

12

11

10

9

Brightness control wheel

A                     B                     C

Figure 7

1. Target  2. field center 3. track 4. track center

1             2               3                 4

7

4. operation of the condenser:

1) the condenser is under the stage, refer to figure 8. The screws on the both sides is for center fixing. The up-and-down hand wheel is for lifting or lowering the collector.

- Adjustment:

a. observe with 40 X , turn down the aperture diaphragm, lower the collector to watch the diaphragm figure.

b. If the collector is not well-adjusted, the diaphragm figure would be unclear and drift the center of the field. Just as figure 9A.

c. Adjust the up-and-down hand wheel to make the edge of the figure clear and the use the two screw to fix the collector in the center, as figure 9B.

d. After all the steps, turn on the aperture diaphragm as a bit bigger than the field, as figure 9C. Lift the colletor until observable.

1        2             3          4              5

6                   7                       8

Figure 8

- Condenser lift hand wheel

- Assitant condenser

- Aperture diaphragm hand wheel

- Condenser lock screw

- Condenser centering hand wheel

- Polarizer

- Field diaphgram ring

- Filter seat

8

A                     B                         C

Figure 9

3                      4                        5

Figure 10

- Observation/ photography lever

- Analyzer plate

- Bertrand centeriing knob

- Compensator

- Analyzer

Figure 11

- Lamp seat

- Lamp

- Lamp centering screw

- Lamp lock screw