MZ61 Series

Stereo Microscope

Operating Manual

Congratulations and Thank You for your purchase of The Microscope MZ61

This precision instrument has been designed to require a minimum of optical and mechanical maintenance. Its excellent design assures years of high quality, reliable service. We recommend you read this entire manual carefully before beginning to use the instrument.

WARNING

PLEASE DISCONNECT PLUG FROM MAINS SOCKET BEFORE REMOVING THE BOTTOM PLATE, OPENING THE LAMP REPLACEMENT DOOR, OR REMOVING THE LAMP HOUSE

WARNING

THE POWER CORD PROVIDED WITH THE EQUIPMENT HAS A GROUNDED PLUG. ALWAYS USE THE POWER CORD WITH A PROPERLY GROUNDED WALL OUTLET.

DO NOT EXPOSE THE INSTRUMENT TO HIGH TEMPERATURES OR HUMIDITY. AVOID USING THE INSTRUMENT IN EXTREMELY DUSTY LOCATIONS

OPERATING TEMPERATURE 5 DEGREE C. TO 35 DEGREE C.

OPERATING HUMIDITY: 20% TO 80% @ 25 DEGREE C.

CAUTION:	NEVER IMMERSE THE INSTRUMENT IN WATER OR SOLVENT

CAUTION:	DO NOT PUT ANY FOREIGN OBJECTS IN THE FRAME OR INTO ANY MOVING MECHANICAL PARTS

Table of Contents

Getting Started	3

System diagram	3

Technical Specifications	4

Components	5

Setup Instructions	7

Basic Operation	7

Replacing Reflected (upper) Light	9

Replacing the Transmitted (lower) Light Lamp	9

Replacing the Fuse	10

Preventative Maintenance	10

Getting Started

This reference guide is written base on whole series, It covers some capable optional attachments and function, you may read the similar operation what you need.

System diagram

Technical Specifications

It has objectives zooming 0.7X to 4.5X

The optical observation tube is inclined to 45°, can be rotated 360°, and the interpupillary distance is adjustable, adjustable range from 55mm to 75 mm.

The left and right observation tube has diopter adjustment ring, adjustable range from +5 to –5 diopter.

Reflected illumination, transmitted illumination or additional illumination can be optional, and their brightness can be adjusted.

It’s frame has arm type and column type to be chosen.

Optical system specification and parameter

| Eyepiece | WF10X | WF15X(optional) | W.D (mm) |
| --- | --- | --- | --- |
| / | Total magnification | Rang of field | Total magnification | Rang of field |  |
|  | 7～45X | Φ28.6～ Ф4.4 | 10.5～67.5X | Ф21.4～Ф3.3 | 88 |
| Additional objective (optional) |
| 0.5X | 3.5～22.5X | Ф57.2～ Ф8.9 | 5.3～33.75X | Ф42.9～ Ф6.7 | 137 |
| 0.75X | 5.3～33.8X | Ф38.1～ Ф5.9 | 7.9～50.6X | Ф28.6～ Ф4.4 | 103 |
| 1X | 7～45X | Φ28.6～ Ф4.4 | 10.5～67.5X | Ф21.4～Ф3.3 | 75 |
| 1.5X | 10.5～67.5X | Ф19～ Ф2.96 | 15.8～101.3X | Ф14.3～ Ф2.2 | 48 |
| 2X | 14～90X | Ф14.3～ Ф2.2 | 21～135X | Ф10.7～ Ф1.7 | 29 |

Components

Setup Instructions

- Remove all parts from their packing materials and retain the packaging in the event you need to transport the product.

Loosen the optical body set screw and install the optical body to the frame by removing its dust cap. Secure optical body with screw. 
DO NOT RELEASE THE OPTICAL BODY FORM YOUR GRASP UNTIL IT IS FIRMLY SECURED TO THE FRAME.

Remove the dust caps from the eyetubes in the Optical Body and insert both eyepieces into the optical body.

Connect the power cord to the instrument and to a suitable power supply.

Basic Operation

- Illumination Controls

As model built-in illumination, Connect the power cord to a suitable power supply, see figure 1. Turn on the power switch, and adjust the transmitted light brightness control. When turn on the reflected light brightness control, reflected light source is switched on, the angle of halogen lamp with the reflector cup can be adjusted.

When use additional illumination, see figure 2. It can be put on the objective tube, you may refer to the illumination user guide before use.

Focusing Controls

Focusing adjustment is accomplished by using the focus knob located comfortably on each side of the frame. This arrangement allows for easy, precise adjustment without drift or discomfort. Turning either of the focus control knobs will raise or lower the optical body.

Interpupillary and Diopter Adjustments, see figure 3.

- Diopter adjustment proper correction for individual vision is accomplished via the diopter adjustment located at both eyepieces, turn the right diopter adjustment ring to “0”, turn the zoom knob to 4.5X, observe at right tube, adjust the focus control knob until image clearly, and then observe at left tube, adjust the left diopter adjustment ring to image clearly, turn the zoom knob to 0.7X, if the image is not clear, don’t adjust zoom knob, observe at right tube, adjust the right diopter adjustment ring to get the image clearly; observe at left tube, adjust the left diopter adjustment ring to get the image clearly ;then turn the zoom knob to 4.5X, If the image isn’t clear, adjust focus control knob to get the image clearly in the same way; Through adjustment can get clearly image from 4.5X to 0.7X.

Proper interpupillary distance, or the distance between eyepieces, is crucial to the comfort of the user. Adjusting the interpupillary distance is accomplished through a “turning” action of the optical head, at figure 3, allowing for quick and easy adjustment.

Additional objective can tighten into the objective tube, See figure 1.

Replacing Reflected (upper) Light

- Disconnect the power plug and wait to be sure the bulb is cool. Loosen the optical body set screw, and remove the optical body form the frame.

Loosen the lamp cover, the halogen bulb with the reflector cup can be removed, Insert the new one into the same fixture. When installing the new bulb, be careful not to touch the glass with your fingers. The new lamp should be supplied in a plastic protective envelope. If not, use a tissue or other medium to grasp the bulb. This will prevent contamination from your hand from reducing the bulb’s intensity and life.

Reinstall the lamp house. Switch on the power, adjustment brightness of illumination, check normal or not.

Replacing the Transmitted (lower) Light Lamp

- Halogen lamp illumination, see figure 4
Disconnect the power plug and wait to be sure the bulb is cool. the halogen lamp socket is located in the base. Remove the slide holder and the table-board. Pull the old lamp straight out of the socket. Insert the new one into the lamp socket. When installing the new bulb, be careful not to touch the glass with your fingers. The new lamp should be supplied in a plastic protective envelope. If not, use a tissue or other medium to grasp the bulb. This will prevent contamination from your hand from reducing the bulb’s intensity and life. Switch on the power, adjustment brightness of illumination, check normal or not.

Fluorescent lamp illumination
Turn off and unplug the microscope. Loosen the screws, see figure 5. remove the old fluorescent lamp from the lamp socket. Insert the new one into the socket. When installing the new bulb, be careful not to touch the glass with your fingers. The new lamp should be supplied in a plastic protective envelope. If not, use a tissue or other medium to grasp the bulb. This will prevent contamination from your hand from reducing the bulb’s intensity and life. Reinstall the chassis. Switch on the power, check normal or not.

Replacing the Fuse

Turn off and unplug the microscope. The fuse socket is located on the back of base, see figure 5. You may unscrew the fuse house and exchange a new fuse, and then screw the fuse house again.

Preventative Maintenance

- Cleaning frame and base

Disconnect the plug from mains socket before cleaning. Clean the frame and base with a soft cloth moistened with a mild detergent solution. Be sure the instrument is dry before using.

Cleaning optical parts

Microscope eyepieces and objectives are coated.

They should not be wiped while dry as dirt or dust may scratch the coating. It is best to remove parts from the frame prior to cleaning. Always blow loose dust away first. Use cotton swabs or lens tissue moistened with a lens cleaner or a small amount of alcohol, then wipe the surface clean with a good quality lens tissue. Solvents such as Xylene should NOT be used as cleaner.

# Text boxes

12

15

14

13

11

10

9

8

7

6

5

4

3

Figure 1

1. Eyepiece  2. Diopter Adjustment Ring  3. Adapter Set Screw  4. Dust Cap  5. Adapter Adjustment Ring  6. Zoom Knob  7. Optical Body Set Screw  8. Halogen lamp with Reflector Cup (with lamp cover)  9. Focus Control Knob  10. Power Switch  11. Reflected Light Brightness Control  12. Transmitted Light Brightness Control  13. Table-board  14. Additional Objective  15. Objective Tube

1

1

Figure 3

2

3

- Diopter Adjustment Ring

- Zoom Knob

2

Figure 2

1. Set Knob  2. Annular Fluorescent Illumination Set Screw  3. Annular Fluorescent Illumination

Digital & Video Camera

1. Lamp Cover  2. Optical Body Set Screw  3. Halogen Lamp with Reflector Cup  4. Power Switch  5. Table-board  6. Halogen Lamp  7. Slide Holder

Figure 4

5

6

7

2

1

4

3

2

1

1. Screw  2. Fluorescent Lamp  3 Lamp Socket  4 Power Switch  5. Fuse House

Figure 5

5

4

3

2

1

Base

Frame

Additional Objective

Additional Illumination

Optical Body

CCD Adapter

Eyepiece